## Slide 1
PRR of FEE and readout electronics (I): Status of PSD electronics development and tests. Readiness to mass production.
Alexander Makhnev on behalf of the PSD team

## Slide 2: PSD electronics architecture
Radiation-hard front-end
  Boards with MPPCs mounted on modules
Radiation-sensitive readout rack
  Readout modules
  HV supply
  Calibration pulse generator
  LV distribution board
Interconnection with 50m signal cables

## Slide 3: Development and tests: Front-end boards
Application: Light readout, temperature sensing
Board includes 10 MPPCs, a temperature sensor and a calibration LED
Light protection components
Improved performance with regard to noise and dark current
No front-end amplification for radiation hardness and high dynamic range
Design: Ready for production
Verification: Done during both cosmic runs at INR at test runs at FAIR
Chart: 
| Dark current, nA per channel
NICA board | 680
FEE board, no light protection | 1000
FEE board, wth light protection | 200
FEE board, with light protection and module cap | 80

## Slide 4
  ADC board was designed for ECAL@PANDA
  LTM9011 ADC;
    125Msps / 80Msps (used now); 14-bit digitization; Selectable Input Ranges: 1VP-P to 2VP-P (used);
  Two FPGA Kintex 7
    32 channels per one FPGA; Separated optical link for each
  Readout rate
    3.2 (4.8) Gb/s; 100 (150) Mb/s per channel
PSD ADC readout board

## Slide 5: Development and tests: Readout module
Application: Signal digitization, HV adjustment, R2D conversion for temperature
Crate-mounted Eurocard-6U module
Assembled from:
  ADC64 board for PANDA experiment
  Interface module (developed at INR)
    Differential ADC interface
    High voltage adjustment
    Input and output zero level adjustment
    Temperature sensor interface
Design: ready for production
Verification: Done during both cosmic runs at INR at test runs at FAIR

## Slide 6: ADC board + interface board
Readout module functional diagram
ADC input channel

## Slide 7: Readout chain noise
Overall noise as measured by ADC under lab conditions is around 0.4 mV RMS
Cosmic ray measurements show clear separation between the 1 MIP signal and the pedestal

## Slide 8: HV adjustment circuit
A common HV is provided to each module
DAC applies a set DC voltage to each channel individually
By adjusting the compensating DAC voltage, MPPC’s voltage can be altered
ADC is decoupled from the line and reads an AC component of the signal
Line remains 50 Ohm terminated with respect to the AC component
Current status:
  Both Readout-side and detector-side PCB with this schematic have been produced
  DAC control firmware is 50% completed

## Slide 9: HV adjustment circuit test
Measured both MPPC linearity vs overvoltage and linearity vs light intensity
Gain can be adjusted via overvoltage alteration
Linearization of the module gain profile is shown to be possible

## Slide 10: Development and tests: calibration pulse generator
Application: driving LEDs for optical calibration
Pulse-time-stable logic pulse generator with remote control
Line drivers for remote application
Design: undergoing reliability tests
Verification: schematics proven to be operation on a evaluation board
Crate-mounted module with the generator is undergoing development

## Slide 11: Development and tests: HV supply
Application: providing HV bias and MPPC current sensing
Remotely controlled HV power supply
  Up to 55V, 10 mA per channel, 48 channels
Per-channel current monitoring
Overcurrent and short-cirquit detection
Temperature and load-stable
Overall supply adjustment with per-channel adjustment done via readout boards
Design: undergoing reliability tests
Verification: schematics proven to be operation on a evaluation board
Crate-mounted module with the supply is undergoing development

## Slide 12: Crate requirements
Readout modules
  Each readout module provides 64 channels
  8 modules needed: 2x6U 280mm air-cooled crates with control panel and power supply (by Wiener), no backplane needed
  Modules are controlled via GBT
  Power is controlled via crate’s interface and/with a custom power distribution board
LED pulser and HV power supply
  Will be implemented in a 3U modules, 1 rack with the same requirements as the above needed
Readout module assembly in crate

## Slide 13: Signal cable verification
Distance between radiation-hard “Detector side” and radiation-sensitive “Readout rack” is up to 60m
10 channel X 46 modules X(50+10spare)m = 27.6 km signal cable
РК50-3-310нг(С)-HF cables and DRAKA CB50 cables were tested
Both cable show acceptable attenuation and skew, transmission of the signals without amplifiers is proven to be possible
For 60 m cable attenuation is 20%
60m cables used during beam tests at mCBM 2021
500 MIPs MPPC signals (HV42)
MPPC pulse waveform with cable’s frequency applied (estimation)
Cable frequency response
MPPC signals propagation via 60m cable

## Slide 14
GBT link is used for control of ADC64 boards
  ADC64 boards send I2C packets with commands for interface boards
    HV adjustment
    Temperature readout
  One of the boards also handles HV supply and pulser control via the same I2C interface
Control scheme
ADC64
(FPGA)
ADC64
(FPGA)

## Slide 15
Architecture: completed
Readout module: ready for production
HV bias generator: schematic tested, ready to production after the end of reliability tests
LED pulser: schematic tested, ready to production after the end of reliability tests
Cabling: verified, ordered
Readout chain in it’s final form is tested on the beam and is operational
Overall status

## Slide 16
No # | Item | CBM | spare (testbenches) | total | To be purchased / manufactured | Purchase / production | comments
1 | ADC board | 8 | 1, (2) | 11 | 9 | Purchase (end of 2022) | 
2 | ADC addon | 8 | 1, (2) | 11 | 11 | Manufactured at INR (middle of 2023) | 
3 | LED pulser board | 1 | 1, (2) | 4 | 4 | Manufactured at INR (middle of 2023) | 
4 | MPPC bias source | 1 | 1, (2) | 4 | 4 | Manufactured at INR (middle of 2023) | 
5 | MPPC board | 46 | 2, (2) | 50 | 50 | Manufactured at INR (middle of 2023) | 
6 | cables | 460 x60m | 0, (20 x60m) | 28,8km | 28,8 km | Purchase (2022 -2023) | 
7 | 6U 280mm Wiener crates | 2 | 0, (0) | 2 | 2 | Purchase (2022 -2023) | 6U 280mm air-cooled crates with control panel and power supply (by Wiener), no backplane needed. For ADC boards
8 | 3U basket crate | 1 | 0, (0) | 1 | 1 | Purchase (2022 -2023) | For LED pulser and MPPC bias source board
Production / purchase summary table of the PSD FEE and readout electronics

## Slide 17
 | 2022 | 2023 Q1 – Q2 | 2023 Q3 – Q4 | 2024 Q1 – Q2
ADC board | Firmware development | purchasing | Quality tests | The Full PSD FEE setup (8 ADCs) installed in crates connected to CRI in GSI.
ADC addon | Development and tests | manufacturing |  | 
MPPC board | manufacturing | Quality tests |  | 
LED pulser board | Development and tests | Quality tests |  | 
MPPC bias source | Development and tests | Quality tests |  | 
cables |  | purchasing |  | 
crates |  |  | purchasing |